## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Fluke |  |  | S/N |  | 343570890
Model Number |  | 8508A |  |  | Test date |  | 24-07-17 17:10:14
Meter Info |  |  |  |  | Ambient Temp. |  | 23.0
Options |  | 8508A |  |  | CERT_AMB?
Firmware |  | 401533-03.05 |  |  | TEMP?
Date of Baseline Cal |  |  |  |  | TEMP? BASE
Date of Certified Cal |  |  |  |  | TEMP? CERTIFIED
CNFTST? |  | 0 |  |  | STANDBY?
DUT Confidence level
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 290524 |  |  | MFC last calibrated |  | 49 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 290524 |  |  | MFC since DCV Zero |  | 49 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529890647135340E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306130764516422E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000045271113747E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000069664541200E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996126071898848E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6059 days 13.72 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 3.13 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.134 | 7.5e-07 |  |  | N/A |  |  | 0.02562 | 1.611e+45
0.0VDC | 0 | -0.11 | 7.5e-07 |  |  | N/A |  |  | 0.017321 | 1.611e+45
0.0VDC | 0 | -1.1 | 7.5e-07 |  |  | N/A |  |  | 0.13944 | 1.611e+45
0.0VDC | 0 | -22 | 7.5e-07 |  |  | N/A |  |  | 2.2361 | 1.611e+45
0.0VDC | 0 | -110 | 7.5e-07 |  |  | N/A |  |  | 12.3603 | 1.611e+45
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 |  |  | 29.5 |  |  | 0.79213 | 315.8
0.02 | 100mV | 0.02 | 44 |  |  | 15.9 |  |  | 0.49081 | 315.8
0.05 | 100mV | 0.05 | 21.5 |  |  | 7.84 |  |  | 0.18077 | 315.8
0.1 | 100mV | 0.1 | 14 | 0.099999 | 0.1 | 4.78 | 3 | 33.385 | 0.069462 | 315.8
-0.01 | 100mV | -0.0099997 | -81.5 |  |  | -28.7 |  |  | -1.5307 | 315.8
-0.02 | 100mV | -0.02 | -44 |  |  | -14.8 |  |  | -0.33167 | 315.8
-0.05 | 100mV | -0.05 | 21.5 |  |  | -5.34 |  |  | -0.20867 | 315.8
-0.1 | 100mV | -0.1 | 14 | -0.099999 | -0.1 | -2.29 | 3 | -15.9941 | -0.095525 | 315.8
0.1 | 1V | 0.1 | 14 |  |  | 1.9 |  |  | 0.23511 | 315.8
0.2 | 1V | 0.2 | 10.25 |  |  | 2.3 |  |  | 0.15092 | 315.8
0.5 | 1V | 0.5 | 5.9 |  |  | 4.62 |  |  | 0.076449 | 315.8
1 | 1V | 1 | 4.7 | 1 | 1 | 4.38 | 1.5 | 88.7797 | 0.035707 | 315.8
-0.1 | 1V | -0.1 | 14 |  |  | 0.3 |  |  | -0.16415 | 315.8
-0.2 | 1V | -0.2 | 10.25 |  |  | 0.9 |  |  | -0.035355 | 315.8
-0.5 | 1V | -0.5 | 5.9 |  |  | 3.64 |  |  | -0.090061 | 315.8
-1 | 1V | -1 | 4.7 | -1 | -1 | 3.62 | 1.5 | 73.375 | -0.026926 | 315.8
1 | 10V | 1 | 4.7 |  |  | 3.3 |  |  | 0.1 | 315.8
2 | 10V | 2 | 4.1 |  |  | 3.85 |  |  | 0.043301 | 315.8
5 | 10V | 5 | 4.1 |  |  | 4.24 |  |  | 0.043333 | 315.8
10 | 10V | 10 | 3.8 | 10 | 10 | 4.38 | 1.5 | 107.2126 | 0.025 | 315.8
19 | 10V | 19.0001 | 3.8158 | 18.9999 | 19.0001 | 6.3789 | 1.5 | 155.5826 | 0.0082287 | 315.8
-1 | 10V | -1 | 4.7 |  |  | 4.2 |  |  | -0.18559 | 315.8
-2 | 10V | -2 | 4.1 |  |  | 4.1 |  |  | -0.096104 | 315.8
-5 | 10V | -5 | 4.1 |  |  | 4.12 |  |  | -0.030184 | 315.8
-10 | 10V | -10 | 3.8 | -10 | -10 | 4.05 | 0.5 | 105.6682 | -0.023511 | 315.8
-19 | 10V | -19.0001 | 3.8158 | -18.9999 | -19.0001 | 5.8947 | 1.5 | 143.7727 | -0.016103 | 315.8
10 | 100V | 10 | 3.8 |  |  | 1.4 |  |  | 0.25221 | 315.8
20 | 100V | 20.0001 | 3.8 |  |  | 4.35 |  |  | 0.082073 | 315.8
50 | 100V | 50.0002 | 7 |  |  | 4.12 |  |  | 0.076011 | 315.8
100 | 100V | 100.0005 | 6 | 99.9994 | 100.0006 | 4.5 | 2 | 71.1512 | 0.033582 | 315.8
-10 | 100V | -10.0001 | 3.8 |  |  | 5.6 |  |  | -0.23979 | 315.8
-20 | 100V | -20.0001 | 3.8 |  |  | 5.95 |  |  | -0.066666 | 315.8
-50 | 100V | -50.0003 | 7 |  |  | 5.36 |  |  | -0.089442 | 315.8
-100 | 100V | -100.0005 | 6 | -99.9994 | -100.0006 | 4.71 | 2 | 74.4716 | -0.041833 | 315.8
-100 | 1000V | -100.0005 | 6 |  |  | 4.8 |  |  | -0.19221 | 315.8
-200 | 1000V | -200.0007 | 5.5 |  |  | 3.55 |  |  | -0.078173 | 315.8
-500 | 1000V | -500.0021 | 8.2 |  |  | 4.22 |  |  | -0.022361 | 315.8
-1000 | 1000V | -1000.0039 | 7.6 | -999.9921 | -1000.0079 | 3.93 | 2 | 50.0079 | -0.021473 | 315.8
100 | 1000V | 100.0003 | 6 |  |  | 3 |  |  | 0.14142 | 315.8
200 | 1000V | 200.0006 | 5.5 |  |  | 3 |  |  | 0.086602 | 315.8
500 | 1000V | 500.0019 | 8.2 |  |  | 3.76 |  |  | 0.03383 | 315.8
1000 | 1000V | 1000.0039 | 7.6 | 999.9921 | 1000.0079 | 3.91 | 2 | 49.7534 | 0.027386 | 315.8
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99961 | 0.99965 | 85 | 0.99953 | 0.99953 | 32.4126 | 20 | 37.1188 | 5.0762 | 315.8
1.9 Ω | 1.9 | 1.9 | 85 | 1.8998 | 1.9001 | 13.4214 | 15 | 15.5496 | 3.1701 | 315.8
10 Ω | 10.0004 | 10.0003 | 26 | 10 | 10.0007 | -4.3998 | 5 | -16.618 | 0.81409 | 315.8
19 Ω | 19.0039 | 19.0038 | 24 | 19.0033 | 19.0044 | -6.3671 | 5 | -25.972 | 0.2723 | 315.8
100 Ω | 100.0053 | 100.0044 | 15 | 100.0035 | 100.0071 | -9.3995 | 3 | -61.4464 | 0.15634 | 315.8
190 Ω | 190.0032 | 190.0018 | 15 | 189.9998 | 190.0067 | -7.5788 | 3 | -49.5443 | 0.06446 | 315.8
1 kΩ | 999.9469 | 999.9365 | 11 | 999.9329 | 999.9609 | -10.4006 | 3 | -91.2189 | 0.11181 | 315.8
1.9 kΩ | 1899.9106 | 1899.8911 | 11 | 1899.884 | 1899.9372 | -10.2636 | 3 | -90.0181 | 0.066809 | 315.8
10 kΩ | 10000.453 | 10000.38 | 9 | 10000.333 | 10000.573 | -7.2997 | 3 | -76.9453 | 0.092793 | 315.8
19 kΩ | 18999.722 | 18999.581 | 9 | 18999.494 | 18999.95 | -7.4212 | 3 | -78.2259 | 0.087282 | 315.8
100 kΩ | 99998.75 | 99998.07 | 11 | 99997.15 | 100000.35 | -6.8001 | 5 | -56.2779 | 0.086604 | 315.8
190 kΩ | 189994.66 | 189993.35 | 11 | 189991.6201 | 189997.6999 | -6.8949 | 5 | -57.0629 | 0.07895 | 315.8
1 MΩ | 999908.7 | 999911.1 | 16 | 999884.7022 | 999932.6978 | 2.4002 | 8 | 13.4176 | 0.18784 | 315.8
1.9 MΩ | 1899855.1 | 1899850.4 | 17 | 1899807.604 | 1899902.596 | -2.4739 | 8 | -13.1671 | 0.038239 | 315.8
10MΩ | 9998564 | 9999497 | 33 | 9998114.065 | 9999013.935 | 93.3134 | 12 | 265.7434 | 0.22362 | 315.8
19 MΩ | 18999500 | 19000190 | 43 | 18998455.03 | 19000544.97 | 36.3167 | 12 | 81.3492 | 0.29066 | 315.8
100MΩ(2W) | 100009390 | 100120520 | 110 | 99980387.28 | 100038392.7 | 1111.1957 | 180 | 526.7573 | 1.6858 | 315.8
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 3e-05 | 5e-05 |  |  | N/A |  |  | 4.3301e-06 | 1.611e+39
0 Ω | 100 Ω | 5e-05 | 5e-05 |  |  | N/A |  |  | 7.8174e-06 | 1.611e+39
0 Ω | 1 KΩ | 0.0004 | 5e-05 |  |  | N/A |  |  | 7.0711e-05 | 1.611e+39
0 Ω | 10 KΩ | 0.004 | 5e-05 |  |  | N/A |  |  | 0.0012019 | 1.611e+39
0 Ω | 100 KΩ | -0.01 | 5e-05 |  |  | N/A |  |  | 0.0052705 | 1.611e+39
0 Ω | 1 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.086603 | 1.611e+39
0 Ω | 10 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.66667 | 1.611e+39
0 Ω | 100 MΩ(2W) | -10 | 5e-05 |  |  | N/A |  |  | 5 | 1.611e+39
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.44762 | 5e-05 |  |  | N/A |  |  | 8.5958e-05 | 1.611e+39
0 Ω | 100 Ω | 0.44735 | 5e-05 |  |  | N/A |  |  | 4.3589e-05 | 1.611e+39
0 Ω | 1 KΩ | 0.4482 | 5e-05 |  |  | N/A |  |  | 8.6603e-05 | 1.611e+39
0 Ω | 10 KΩ | 0.458 | 5e-05 |  |  | N/A |  |  | 0.00086603 | 1.611e+39
0 Ω | 100 KΩ | 0.45 | 5e-05 |  |  | N/A |  |  | 0.005 | 1.611e+39
0 Ω | 1 MΩ | 0.4 | 5e-05 |  |  | N/A |  |  | 0.088192 | 1.611e+39
0 Ω | 10 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.66667 | 1.611e+39
0 Ω | 100 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 7.0711 | 1.611e+39
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99997 | 77 | 0.99992 | 1.0001 | -30 | 10 | -38.6366 | 18.0283 | 315.8
2VAC/1kHz | 10 VAC | 1.9999 | 73.5 | 1.9999 | 2.011 | -25 | 10 | -33.7031 | 6.1803 | 315.8
5VAC/1kHz | 10 VAC | 4.9999 | 84 | 4.9996 | 5.0004 | -12 | 10 | -14.1855 | 3.5746 | 315.8
10VAC/1kHz | 10 VAC | 9.9999 | 77 | 9.9992 | 1.006 | -9 | 10 | -11.591 | 1.9221 | 315.8
12VAC/1kHz | 10 VAC | 11.9999 | 75.833 | 11.9991 | 1.0001 | -9.1667 | 10 | -11.9842 | 1.9593 | 315.8
15VAC/1kHz | 10 VAC | 14.9999 | 74.667 | 14.9989 | 1.0057 | -9.3333 | 10 | -12.3893 | 1.5072 | 315.8
19VAC/1kHz | 10 VAC | 18.9998 | 73.684 | 18.9986 | 19.0014 | -10.5263 | 10 | -14.156 | 1.0116 | 315.8
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009971 | 5500 | 0.0009945 | 0.0010055 | -2900 | 20 | -52.7269 | 150.4363 | 315.8
0.001 V, 20 Hz |  | 0.000998 | 5200 | 0.0009948 | 0.0010052 | -2000 | 20 | -38.4613 | 142.6854 | 315.8
0.001 V, 30 Hz |  | 0.0009981 | 5200 | 0.0009948 | 0.0010052 | -1900 | 20 | -36.5382 | 33.3968 | 315.8
0.001 V, 40 Hz |  | 0.0009982 | 5100 | 0.0009949 | 0.0010051 | -1800 | 20 | -35.2938 | 116.877 | 315.8
0.001 V, 55 Hz |  | 0.0009982 | 5100 | 0.0009949 | 0.0010051 | -1800 | 20 | -35.2938 | 92.9634 | 315.8
0.001 V, 300 Hz |  | 0.0009984 | 5100 | 0.0009949 | 0.0010051 | -1600 | 20 | -31.3723 | 364.5888 | 315.8
0.001 V, 1 KHz |  | 0.0009979 | 5100 | 0.0009949 | 0.0010051 | -2100 | 20 | -41.1762 | 307.9649 | 315.8
0.001 V, 10 KHz |  | 0.0009984 | 5100 | 0.0009949 | 0.0010051 | -1600 | 20 | -31.3723 | 187.3827 | 315.8
0.001 V, 20 KHz |  | 0.0009986 | 5100 | 0.0009949 | 0.0010051 | -1400 | 20 | -27.4508 | 105.557 | 315.8
0.001 V, 30 KHz |  | 0.0009979 | 5340 | 0.00099466 | 0.0010053 | -2100 | 20 | -39.3256 | 109.5206 | 315.8
0.001 V, 50 KHz |  | 0.0009959 | 5340 | 0.00099466 | 0.0010053 | -4100 | 20 | -76.7785 | 101.7967 | 315.8
0.001 V, 100 KHz |  | 0.0009818 | 8800 | 0.0009912 | 0.0010088 | -18200 | 50 | -206.8148 | 79.6227 | 315.8
0.001 V, 200 KHz |  | 0.0009275 | 16100 | 0.0009839 | 0.0010161 | -72500 | 100 | -450.3019 | 107.8167 | 315.8
0.001 V, 300 KHz |  | 0.0008949 | 16100 | 0.0009839 | 0.0010161 | -105100 | 100 | -652.7824 | 147.8239 | 315.8
0.001 V, 500 KHz |  | 0.0008571 | 31500 | 0.0009685 | 0.0010315 | -142900 | 200 | -453.6417 | 182.4142 | 315.8
0.001 V, 1 MHz |  | 0.0007719 | 33000 | 0.000967 | 0.001033 | -228100 | 300 | -691.1836 | 262.6749 | 315.8
0.01 V, 10 Hz |  | 0.0099871 | 1100 | 0.009989 | 0.010011 | -1290 | 20 | -117.2533 | 129.3093 | 315.8
0.01 V, 20 Hz |  | 0.0099961 | 800 | 0.009992 | 0.010008 | -390 | 20 | -48.7348 | 11.3083 | 315.8
0.01 V, 30 Hz |  | 0.0099974 | 800 | 0.009992 | 0.010008 | -260 | 20 | -32.4898 | 11.3068 | 315.8
0.01 V, 40 Hz |  | 0.0099981 | 700 | 0.009993 | 0.010007 | -190 | 20 | -27.1318 | 6.0104 | 315.8
0.01 V, 55 Hz |  | 0.0099984 | 700 | 0.009993 | 0.010007 | -160 | 20 | -22.8478 | 11.3057 | 315.8
0.01 V, 300 Hz |  | 0.0099988 | 700 | 0.009993 | 0.010007 | -120 | 20 | -17.1359 | 28.6294 | 315.8
0.01 V, 1 KHz |  | 0.0099988 | 700 | 0.009993 | 0.010007 | -120 | 20 | -17.1359 | 33.5864 | 315.8
0.01 V, 10 KHz |  | 0.0099999 | 700 | 0.009993 | 0.010007 | -10 | 20 | -1.428 | 16.5833 | 315.8
0.01 V, 20 KHz |  | 0.01 | 700 | 0.009993 | 0.010007 | 0 | 20 | 0 | 15.0923 | 315.8
0.01 V, 30 KHz |  | 0.010001 | 940 | 0.0099906 | 0.010009 | 100 | 20 | 10.6359 | 8.8183 | 315.8
0.01 V, 50 KHz |  | 0.010003 | 940 | 0.0099906 | 0.010009 | 340 | 20 | 36.162 | 4.9983 | 315.8
0.01 V, 100 KHz |  | 0.010009 | 1600 | 0.009984 | 0.010016 | 870 | 50 | 54.3485 | 10.5318 | 315.8
0.01 V, 200 KHz |  | 0.0099962 | 2600 | 0.009974 | 0.010026 | -380 | 100 | -14.6046 | 4.4113 | 315.8
0.01 V, 300 KHz |  | 0.009993 | 2600 | 0.009974 | 0.010026 | -700 | 100 | -26.9032 | 9.7251 | 315.8
0.01 V, 500 KHz |  | 0.0099884 | 4500 | 0.009955 | 0.010045 | -1160 | 200 | -25.7524 | 7.8264 | 315.8
0.01 V, 1 MHz |  | 0.0097837 | 6000 | 0.0099399 | 0.01006 | -21630 | 300 | -360.0502 | 5.1105 | 315.8
0.1 V, 10 Hz |  | 0.099872 | 660 | 0.099934 | 0.10007 | -1280 | 20 | -193.8504 | 145.7161 | 315.8
0.1 V, 20 Hz |  | 0.099967 | 300 | 0.09997 | 0.10003 | -330 | 20 | -109.7564 | 4.3826 | 315.8
0.1 V, 30 Hz |  | 0.099982 | 300 | 0.09997 | 0.10003 | -182 | 20 | -60.5323 | 2.2427 | 315.8
0.1 V, 40 Hz |  | 0.099987 | 195 | 0.09998 | 0.10002 | -131 | 20 | -66.8289 | 1.7162 | 315.8
0.1 V, 55 Hz |  | 0.09999 | 195 | 0.09998 | 0.10002 | -102 | 20 | -52.0347 | 1.9223 | 315.8
0.1 V, 300 Hz |  | 0.099993 | 195 | 0.09998 | 0.10002 | -71 | 20 | -36.2202 | 2.2424 | 315.8
0.1 V, 1 KHz |  | 0.099992 | 195 | 0.09998 | 0.10002 | -75 | 20 | -38.2608 | 2.7133 | 315.8
0.1 V, 10 KHz |  | 0.099992 | 195 | 0.09998 | 0.10002 | -79 | 20 | -40.3014 | 1.8561 | 315.8
0.1 V, 20 KHz |  | 0.099987 | 195 | 0.09998 | 0.10002 | -133 | 20 | -67.8492 | 1.2695 | 315.8
0.1 V, 30 KHz |  | 0.099983 | 400 | 0.09996 | 0.10004 | -166 | 20 | -41.4482 | 1.3336 | 315.8
0.1 V, 50 KHz |  | 0.099984 | 400 | 0.09996 | 0.10004 | -163 | 20 | -40.6992 | 1.4816 | 315.8
0.1 V, 100 KHz |  | 0.099982 | 1050 | 0.099895 | 0.10011 | -185 | 50 | -17.5991 | 2.3984 | 315.8
0.1 V, 200 KHz |  | 0.09996 | 1240 | 0.099876 | 0.10012 | -398 | 100 | -31.9929 | 0.88227 | 315.8
0.1 V, 300 KHz |  | 0.099936 | 1240 | 0.099876 | 0.10012 | -635 | 100 | -51.044 | 1.2026 | 315.8
0.1 V, 500 KHz |  | 0.099906 | 1900 | 0.099809 | 0.10019 | -941 | 200 | -49.2542 | 1.1315 | 315.8
0.1 V, 1 MHz |  | 0.099645 | 4000 | 0.099599 | 0.1004 | -3547 | 300 | -88.4266 | 2.0689 | 315.8
1 V, 10 Hz |  | 0.99873 | 600 | 0.9994 | 1.0006 | -1268 | 10 | -211.304 | 125.1311 | 315.8
1 V, 20 Hz |  | 0.99972 | 180 | 0.99982 | 1.0002 | -277 | 10 | -153.652 | 3.4651 | 315.8
1 V, 30 Hz |  | 0.99987 | 180 | 0.99982 | 1.0002 | -131 | 10 | -72.6657 | 3.7901 | 315.8
1 V, 40 Hz |  | 0.99992 | 77 | 0.99992 | 1.0001 | -77 | 10 | -99.1672 | 1.1668 | 315.8
1 V, 55 Hz |  | 0.99995 | 77 | 0.99992 | 1.0001 | -49 | 10 | -63.1064 | 1.8029 | 315.8
1 V, 300 Hz |  | 0.99999 | 77 | 0.99992 | 1.0001 | -15 | 10 | -19.3183 | 1.2693 | 315.8
1 V, 1 KHz |  | 0.99998 | 77 | 0.99992 | 1.0001 | -17 | 10 | -21.8941 | 2.4552 | 315.8
1 V, 10 KHz |  | 0.99999 | 77 | 0.99992 | 1.0001 | -14 | 10 | -18.0304 | 0.92797 | 315.8
1 V, 20 KHz |  | 0.99998 | 77 | 0.99992 | 1.0001 | -24 | 10 | -30.9093 | 1.1181 | 315.8
1 V, 30 KHz |  | 0.99998 | 140 | 0.99986 | 1.0001 | -21 | 10 | -14.9619 | 0.7265 | 315.8
1 V, 50 KHz |  | 1 | 140 | 0.99986 | 1.0001 | 10 | 20 | 7.0711 | 1.3229 | 315.8
1 V, 100 KHz |  | 1.0001 | 310 | 0.99969 | 1.0003 | 74 | 30 | 23.76 | 1.2359 | 315.8
1 V, 200 KHz |  | 1.0001 | 550 | 0.99945 | 1.0006 | 84 | 70 | 15.1505 | 1.1179 | 315.8
1 V, 300 KHz |  | 1.0001 | 550 | 0.99945 | 1.0006 | 142 | 70 | 25.6116 | 0.88179 | 315.8
1 V, 500 KHz |  | 1.0009 | 1400 | 0.9986 | 1.0014 | 945 | 100 | 67.3285 | 0.88108 | 315.8
1 V, 1 MHz |  | 1.0076 | 3000 | 0.99699 | 1.003 | 7597 | 200 | 252.6725 | 4.9236 | 315.8
10 V, 10 Hz |  | 9.9872 | 600 | 9.994 | 10.006 | -1275 | 10 | -212.4705 | 142.5487 | 315.8
10 V, 20 Hz |  | 9.9974 | 180 | 9.9982 | 10.0018 | -264 | 10 | -146.4409 | 3.6903 | 315.8
10 V, 30 Hz |  | 9.9988 | 180 | 9.9982 | 10.0018 | -115 | 10 | -63.7905 | 1.7323 | 315.8
10 V, 40 Hz |  | 9.9994 | 77 | 9.9992 | 10.0008 | -64 | 10 | -82.4247 | 1.6416 | 315.8
10 V, 55 Hz |  | 9.9997 | 77 | 9.9992 | 10.0008 | -35 | 10 | -45.076 | 1.9437 | 315.8
10 V, 300 Hz |  | 10 | 77 | 9.9992 | 10.0008 | 1 | 10 | 1.2879 | 1.6159 | 315.8
10 V, 1 KHz |  | 9.9999 | 77 | 9.9992 | 10.0008 | -9 | 10 | -11.591 | 2.6926 | 315.8
10 V, 10 KHz |  | 10.0003 | 77 | 9.9992 | 10.0008 | 28 | 10 | 36.0608 | 1.3333 | 315.8
10 V, 20 KHz |  | 10.0009 | 77 | 9.9992 | 10.0008 | 95 | 10 | 122.3492 | 1.054 | 315.8
10 V, 30 KHz |  | 10.0016 | 140 | 9.9986 | 10.0014 | 164 | 10 | 116.8452 | 1.5363 | 315.8
10 V, 50 KHz |  | 10.0025 | 140 | 9.9986 | 10.0014 | 246 | 20 | 173.9483 | 0.70693 | 315.8
10 V, 100 KHz |  | 10.0027 | 270 | 9.9973 | 10.0027 | 271 | 30 | 99.7565 | 1.3225 | 315.8
10 V, 200 KHz |  | 10.0067 | 670 | 9.9933 | 10.0067 | 667 | 70 | 99.0133 | 1.2684 | 315.8
10 V, 300 KHz |  | 10.0115 | 670 | 9.9933 | 10.0067 | 1153 | 70 | 171.1579 | 1.3002 | 315.8
10 V, 500 KHz |  | 10.0587 | 1700 | 9.983 | 10.017 | 5869 | 100 | 344.6395 | 0.86097 | 315.8
10 V, 1 MHz |  | 9.9944 | 3500 | 9.9649 | 10.0351 | -557 | 200 | -15.8884 | 5.7357 | 315.8
19 V, 10 Hz |  | 18.9812 | 552.63 | 18.9895 | 19.0105 | -986.8421 | 10 | -178.5427 | 130.7686 | 315.8
19 V, 20 Hz |  | 18.9961 | 165.79 | 18.9968 | 19.0032 | -203.1579 | 10 | -122.317 | 4.3045 | 315.8
19 V, 30 Hz |  | 18.9983 | 165.79 | 18.9968 | 19.0032 | -86.8421 | 10 | -52.2858 | 1.8692 | 315.8
19 V, 40 Hz |  | 18.9991 | 73.684 | 18.9986 | 19.0014 | -46.3158 | 10 | -62.2863 | 1.8876 | 315.8
19 V, 55 Hz |  | 18.9995 | 73.684 | 18.9986 | 19.0014 | -24.7368 | 10 | -33.2666 | 1.1096 | 315.8
19 V, 300 Hz |  | 19.0001 | 73.684 | 18.9986 | 19.0014 | 2.6316 | 10 | 3.539 | 0.9116 | 315.8
19 V, 1 KHz |  | 18.9998 | 73.684 | 18.9986 | 19.0014 | -8.4211 | 10 | -11.3248 | 0.94069 | 315.8
19 V, 10 KHz |  | 19.0005 | 73.684 | 18.9986 | 19.0014 | 26.3158 | 10 | 35.39 | 1.147 | 315.8
19 V, 20 KHz |  | 19.0017 | 73.684 | 18.9986 | 19.0014 | 87.3684 | 10 | 117.4947 | 1.099 | 315.8
19 V, 30 KHz |  | 19.0029 | 130.53 | 18.9975 | 19.0025 | 153.6842 | 10 | 117.3946 | 1.1503 | 315.8
19 V, 50 KHz |  | 19.0043 | 130.53 | 18.9975 | 19.0025 | 225.7895 | 20 | 170.9836 | 1.2089 | 315.8
19 V, 100 KHz |  | 19.0041 | 251.05 | 18.9952 | 19.0048 | 216.8421 | 30 | 85.7639 | 1.1162 | 315.8
19 V, 200 KHz |  | 19.01 | 589.47 | 18.9887 | 19.0113 | 525.7895 | 70 | 88.5746 | 0.83635 | 315.8
19 V, 300 KHz |  | 19.0167 | 589.47 | 18.9887 | 19.0113 | 877.3684 | 70 | 147.8017 | 0.68451 | 315.8
19 V, 500 KHz |  | 19.1004 | 1463.2 | 18.9721 | 19.0279 | 5284.7368 | 100 | 360.3361 | 1.2278 | 315.8
19 V, 1 MHz |  | 18.9527 | 3073.7 | 18.9415 | 19.0585 | -2491.5789 | 200 | -80.8902 | 5.1224 | 315.8
100 V, 10 Hz |  | 99.8759 | 600 | 99.94 | 100.06 | -1241 | 10 | -206.8046 | 118.0696 | 315.8
100 V, 20 Hz |  | 99.9746 | 180 | 99.982 | 100.018 | -254 | 10 | -140.8938 | 4.7709 | 315.8
100 V, 30 Hz |  | 99.9895 | 180 | 99.982 | 100.018 | -105 | 10 | -58.2435 | 3.0003 | 315.8
100 V, 40 Hz |  | 99.9942 | 85 | 99.9914 | 100.0086 | -58 | 10 | -67.7679 | 2.9061 | 315.8
100 V, 55 Hz |  | 99.9975 | 85 | 99.9914 | 100.0086 | -25 | 10 | -29.2103 | 2.5496 | 315.8
100 V, 300 Hz |  | 100.0007 | 85 | 99.9914 | 100.0086 | 7 | 10 | 8.1789 | 1.5899 | 315.8
100 V, 1 KHz |  | 100.0001 | 85 | 99.9914 | 100.0086 | 1 | 10 | 1.1684 | 1.8105 | 315.8
100 V, 10 KHz |  | 100.0023 | 85 | 99.9914 | 100.0086 | 23 | 10 | 26.8735 | 1.4142 | 315.8
100 V, 20 KHz |  | 100.0103 | 85 | 99.9914 | 100.0086 | 103 | 10 | 120.3465 | 0.86594 | 315.8
100 V, 30 KHz |  | 100.018 | 240 | 99.976 | 100.024 | 180 | 10 | 74.935 | 1.4527 | 315.8
100 V, 50 KHz |  | 100.0266 | 240 | 99.9759 | 100.0241 | 266 | 20 | 110.4505 | 1.6579 | 315.8
100 V, 100 KHz |  | 100.0332 | 600 | 99.9399 | 100.0601 | 332 | 30 | 55.2643 | 6.739 | 315.8
100 V, 200 KHz |  | 100.1198 | 2600 | 99.74 | 100.26 | 1198 | 50 | 46.0684 | 14.3712 | 315.8
1000 V, 55 Hz |  | 999.962 | 79 | 999.9196 | 1000.0804 | -38 | 15 | -47.257 | 2.0481 | 315.8
1000 V, 300 Hz |  | 1000.009 | 79 | 999.9196 | 1000.0804 | 9 | 15 | 11.1924 | 1.0541 | 315.8
700 V, 1 KHz |  | 700.012 | 80.714 | 699.9335 | 700.0665 | 17.1429 | 50 | 18.0554 | 1.8132 | 315.8
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.9993e-07 | 10045.653 |  |  | -65 |  |  | 9.5749 | 315.8
10 µADC | 1e-05 | 1.0001e-05 | 1044.9467 |  |  | 51 |  |  | 1.732 | 315.8
50 µADC | 5e-05 | 5.0003e-05 | 245 |  |  | 62.6 |  |  | 0.58875 | 315.8
100 µADC | 0.0001 | 0.00010001 | 145 | 9.9985e-05 | 0.00010001 | 63.65 | 7 | 43.8455 | 0.17077 | 315.8
-100 µADC | -0.0001 | -0.00010001 | 145 | -9.9985e-05 | -0.00010001 | 66 | 7 | 45.4643 | -0.17319 | 315.8
-50 µADC | -5e-05 | -5.0003e-05 | 245 |  |  | 66.7 |  |  | -0.48987 | 315.8
-10 µADC | -1e-05 | -1.0001e-05 | -1044.9206 |  |  | 76 |  |  | -1.2582 | 315.8
0.1 mADC | 0.0001 | 0.00010001 | 145 |  |  | 63.5 |  |  | 1.2909 | 315.8
0.5 mADC | 0.0005 | 0.00050002 | 65 |  |  | 45.6 |  |  | 0 | 315.8
1 mADC | 0.001 | 0.001 | 55 | 0.00099994 | 0.0010001 | 46.3 | 7 | 83.5082 | 0.081646 | 315.8
-1 mADC | -0.001 | -0.001 | 55 | -0.00099994 | -0.0010001 | 46.9 | 7 | 84.5904 | -0.18257 | 315.8
-0.5 mADC | -0.0005 | -0.00050002 | 65 |  |  | 47.3 |  |  | -0.19148 | 315.8
1 mADC | 0.001 | 0.001 | 55 |  |  | 44 |  |  | 0.49998 | 315.8
5 mADC | 0.005 | 0.0050002 | 65 |  |  | 40.3 |  |  | 0.47608 | 315.8
10 mADC | 0.01 | 0.01 | 55 | 0.0099994 | 0.010001 | 40.95 | 7 | 73.8588 | 0.23804 | 315.8
-10 mADC | -0.01 | -0.01 | 55 | -0.0099994 | -0.010001 | 40.95 | 7 | 73.8588 | -0.1732 | 315.8
-5 mADC | -0.005 | -0.0050002 | 65 |  |  | 41.6 |  |  | -0.23093 | 315.8
10 mADC | 0.01 | 0.01 | 55 |  |  | 28 |  |  | 0.81647 | 315.8
50 mADC | 0.05 | 0.050004 | 75 |  |  | 89.8 |  | 119.7333 | 0.19998 | 315.8
100 mADC | 0.1 | 0.10001 | 65 | 0.099993 | 0.10001 | 95.5 | 7 | 146.0784 | 0.20614 | 315.8
-100 mADC | -0.1 | -0.10001 | 65 | -0.099993 | -0.10001 | 107.65 | 7 | 164.6633 | -0.20613 | 315.8
-50 mADC | -0.05 | -0.050006 | 75 |  |  | 113 |  |  | -0.32656 | 315.8
0.1 ADC | 0.1 | 0.099941 | 65 |  |  | -590.5 |  |  | 0.57769 | 315.8
0.5 ADC | 0.5 | 0.49981 | 135 |  |  | -376.3 |  |  | 0.19156 | 315.8
-0.5 ADC | -0.5 | -0.49989 | 135 |  |  | -221.1 |  |  | -0.34164 | 315.8
-1.0 ADC | -1 | -0.99973 | 105 | -0.99989 | -1.0001 | -269.95 | 20 | -252.5546 | -0.095769 | 315.8
1.0 ADC | 1 | 0.99965 | 105 | 0.99989 | 1.0001 | -346.75 | 20 | -324.4056 | 0.2161 | 315.8
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.991e-06 | 3653.288 |  |  | -900 |  |  | 71.2446 | 315.8
10 µA, 20 Hz | 1e-05 | 9.9931e-06 | 2851.9536 |  |  | -685 |  |  | 9.5808 | 315.8
10 µA, 30 Hz | 1e-05 | 9.9932e-06 | 2851.9393 |  |  | -680 |  |  | 8.1705 | 315.8
10 µA, 40 Hz | 1e-05 | 9.9925e-06 | 2121.5806 |  |  | -745 |  |  | 9.5814 | 315.8
10 µA 55 Hz | 1e-05 | 9.9928e-06 | 2121.5381 |  |  | -725 |  |  | 20.6305 | 315.8
10 µA, 300 Hz | 1e-05 | 9.993e-06 | 2121.485 |  |  | -40 |  |  | 5.0035 | 315.8
10 µA, 1 KHz | 1e-05 | 9.9924e-06 | 2121.6018 |  |  | -755 |  |  | 9.5815 | 315.8
10 µA, 5 KHz | 1e-05 | 9.9831e-06 | 5509.2831 |  |  | -1685 |  |  | 9.5904 | 315.8
10 µA, 10 KHz | 1e-05 | 9.9541e-06 | 11553.0864 |  |  | -4595 |  |  | 9.6185 | 315.8
100 µA, 10 Hz | 0.0001 | 9.9998e-05 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 | -4.2047 | 81.1329 | 315.8
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 | 9.994e-05 | 0.00010006 | 99.5 | 50 | 16.5261 | 2.6297 | 315.8
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 | 91 | 50 | 15.1143 | 1.8256 | 315.8
100 µA, 40 Hz | 0.0001 | 0.0001 | 320 | 9.9968e-05 | 0.00010003 | 11.5 | 50 | 3.5507 | 1.7078 | 315.8
100 µA 55 Hz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 | -12.3502 | 2.9999 | 315.8
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 | 9.9968e-05 | 0.00010003 | 67.5 | 50 | 20.8409 | 0.95736 | 315.8
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 | 21 | 50 | 6.4838 | 1.5 | 315.8
100 µA, 5 KHz | 0.0001 | 9.9915e-05 | 0.1 | 9.999e-05 | 0.00010001 | -850 | 100 | -849.9996 | 1.2594 | 315.8
100 µA, 10 KHz | 0.0001 | 9.9653e-05 | 0.25 | 9.9939e-05 | 0.00010006 | -3468 | 300 | -1155.9996 | 1.7381 | 315.8
1mA, 10 Hz | 0.001 | 0.00099993 | 700 | 0.0009993 | 0.0010007 | -75 | 40 | -10.6968 | 123.9054 | 315.8
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 | 57 | 40 | 14.5391 | 2.6456 | 315.8
1 mA, 30 Hz | 0.001 | 0.001 | 390 | 0.00099961 | 0.0010004 | 47.5 | 40 | 12.1159 | 1.4141 | 315.8
1 mA, 40 Hz | 0.001 | 0.00099997 | 160 | 9.9939e-05 | 0.00010006 | -32 | 40 | -19.4029 | 1.2583 | 315.8
1 mA, 55 Hz | 0.001 | 0.001 | 160 | 0.00099984 | 0.0010002 | -4.5 | 40 | -2.7285 | 2.3805 | 315.8
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 | 27 | 40 | 16.3712 | 0.49999 | 315.8
1 mA, 1 KHz | 0.001 | 0.001 | 160 | 0.00099984 | 0.0010002 | 14 | 40 | 8.4887 | 1.1547 | 315.8
1 mA, 5 KHz | 0.001 | 0.00099996 | 1000 | 9.9939e-05 | 0.00010006 | -36 | 70 | -3.5912 | 1.9149 | 315.8
1 mA, 10 KHz | 0.001 | 0.00099974 | 0.25 | 0.0009998 | 0.0010002 | -265 | 200 | -132.4999 | 0.50013 | 315.8
10 mA, 10 Hz | 0.01 | 0.0099992 | 700 | 9.9939e-05 | 0.00010006 | -76.5 | 40 | -10.9108 | 175.6921 | 315.8
10 mA, 20 Hz | 0.01 | 0.010001 | 390 | 0.0099961 | 0.010004 | 56.5 | 40 | 14.4116 | 3.8728 | 315.8
10 mA, 30 Hz | 0.01 | 0.01 | 390 | 9.9939e-05 | 0.00010006 | 49.5 | 40 | 12.6261 | 1.2909 | 315.8
10 mA, 40 Hz | 0.01 | 0.0099997 | 160 | 0.0099984 | 0.010002 | -33 | 40 | -20.0092 | 1.7321 | 315.8
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 | -2.5 | 40 | -1.5158 | 1.893 | 315.8
10 mA, 300 Hz | 0.01 | 0.01 | 160 | 0.0099984 | 0.010002 | 30 | 40 | 18.1902 | 0.81647 | 315.8
10 mA, 1 KHz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 | 21 | 40 | 12.7331 | 1.5 | 315.8
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 | 0.00999 | 0.01001 | 74 | 70 | 7.3819 | 1.1546 | 315.8
10 mA, 10 KHz | 0.01 | 0.010001 | 0.25 | 9.9939e-05 | 0.00010006 | 115 | 200 | 57.5 | 0.99989 | 315.8
100 mA, 10 Hz | 0.1 | 0.099994 | 700 | 0.09993 | 0.10007 | -63 | 40 | -8.9853 | 134.2334 | 315.8
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 | 112 | 40 | 28.5681 | 1.2582 | 315.8
100 mA, 30 Hz | 0.1 | 0.10001 | 390 | 0.099961 | 0.10004 | 106 | 40 | 27.0376 | 2.6297 | 315.8
100 mA, 40 Hz | 0.1 | 0.1 | 160 | 9.9939e-05 | 0.00010006 | 19 | 40 | 11.5204 | 0.49999 | 315.8
100 mA, 55 Hz | 0.1 | 0.1 | 160 | 0.099984 | 0.10002 | 48 | 40 | 29.1043 | 2.0615 | 315.8
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 | 83.5 | 40 | 50.6293 | 3.162 | 315.8
100 mA, 1 KHz | 0.1 | 0.10001 | 160 | 0.099984 | 0.10002 | 79 | 40 | 47.9008 | 4.425 | 315.8
100 mA, 5 KHz | 0.1 | 0.10002 | 0.1 | 9.9939e-05 | 0.00010006 | 242 | 70 | 345.7139 | 2.2168 | 315.8
100 mA, 10 KHz | 0.1 | 0.10005 | 0.25 | 0.09998 | 0.10002 | 518.5 | 200 | 259.2498 | 2.8853 | 315.8
1 A, 10 Hz | 1 | 0.99949 | -1 | 9.9939e-05 | 0.00010006 | -511 | 40 | -1277.101 | 141.5656 | 315.8
1 A, 20 Hz | 1 | 0.9997 | -1 | 0.99996 | 1 | -296 | 40 | -739.7689 | 4.436 | 315.8
1 A, 30 Hz | 1 | 0.99969 | -1 | 9.9939e-05 | 0.00010006 | -305.5 | 40 | -763.5114 | 1.7083 | 315.8
1 A, 40 Hz | 1 | 0.99961 | 640 | 0.99936 | 1.0006 | -392 | 40 | -61.1307 | 4.6923 | 315.8
1 A, 55 Hz | 1 | 0.99964 | 640 | 9.9939e-05 | 0.00010006 | -361.5 | 40 | -56.3744 | 4.0429 | 315.8
1 A, 300 Hz | 1 | 0.99968 | 640 | 0.99936 | 1.0006 | -324 | 40 | -50.5264 | 3.001 | 315.8
1 A, 1 KHz | 1 | 0.99972 | 640 | 9.9939e-05 | 0.00010006 | -275.5 | 40 | -42.963 | 3.56 | 315.8
1 A, 5 KHz | 1 | 0.99978 | 800 | 0.9992 | 1.0008 | -216 | 70 | -26.8972 | 0.81667 | 315.8
1 A, 10 KHz | 1 | 0.9995 | 0.82 | 9.9939e-05 | 0.00010006 | -501.5 | 200 | -250.7479 | 1.7087 | 315.8
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.01 | 278.28 | 146.22 | 434.25 | 33.72 | 27.74 |  | 920.22