## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Fluke |  |  | S/N |  | 343570890
Model Number |  | 8508A |  |  | Test date |  | 24-07-22 16:34:47
Meter Info |  |  |  |  | Ambient Temp. |  | 23.0
Options |  | 8508A |  |  | CERT_AMB?
Firmware |  | 401533-03.05 |  |  | TEMP?
Date of Baseline Cal |  |  |  |  | TEMP? BASE
Date of Certified Cal |  |  |  |  | TEMP? CERTIFIED
CNFTST? |  | 0 |  |  | STANDBY?
DUT Confidence level
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 290524 |  |  | MFC last calibrated |  | 54 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 290524 |  |  | MFC since DCV Zero |  | 54 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529890647135340E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306130764516422E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000045271113747E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000069664541200E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996126071898848E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6064 days 12.95 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 7.37 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.553 | 7.5e-07 |  |  | N/A |  |  | 0.17905 | 1.611e+45
0.0VDC | 0 | -0.06 | 7.5e-07 |  |  | N/A |  |  | 0.056125 | 1.611e+45
0.0VDC | 0 | -1.1 | 7.5e-07 |  |  | N/A |  |  | 0.14142 | 1.611e+45
0.0VDC | 0 | -23 | 7.5e-07 |  |  | N/A |  |  | 2.4495 | 1.611e+45
0.0VDC | 0 | -130 | 7.5e-07 |  |  | N/A |  |  | 12.693 | 1.611e+45
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.010001 | 81.5 | [5 empty] | 1.1147 | 115.8
0.02 | 100mV | 0.020001 | 44 | [5 empty] | 0.31234 | 115.8
0.05 | 100mV | 0.050001 | 21.5 | [5 empty] | 0.19773 | 115.8
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.067102 | 115.8
-0.01 | 100mV | -0.0099989 | -81.5 | [5 empty] | -1.45 | 115.8
-0.02 | 100mV | -0.019999 | -44 | [5 empty] | -0.53609 | 115.8
-0.05 | 100mV | -0.049999 | 21.5 | [5 empty] | -0.1261 | 115.8
-0.1 | 100mV | -0.099999 | 14 |  |  |  | 3 |  | -0.051344 | 115.8
0.1 | 1V | 0.1 | 14 | [5 empty] | 0.1236 | 115.8
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.14386 | 115.8
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.079372 | 115.8
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.028723 | 115.8
-0.1 | 1V | -0.1 | 14 | [5 empty] | -0.13333 | 115.8
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.097183 | 115.8
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.12845 | 115.8
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.045674 | 115.8
1 | 10V | 1 | 4.7 | [5 empty] | 0.13944 | 115.8
2 | 10V | 2 | 4.1 | [5 empty] | 0.075 | 115.8
5 | 10V | 5 | 4.1 | [5 empty] | 0.037416 | 115.8
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.016159 | 115.8
19 | 10V | 19.0001 | 3.8158 |  |  |  | 1.5 |  | 0.024325 | 115.8
-1 | 10V | -1 | 4.7 | [5 empty] | -0.17638 | 115.8
-2 | 10V | -2 | 4.1 | [5 empty] | -0.080794 | 115.8
-5 | 10V | -5 | 4.1 | [5 empty] | -0.031623 | 115.8
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.031269 | 115.8
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.022295 | 115.8
10 | 100V | 10 | 3.8 | [5 empty] | 0.25055 | 115.8
20 | 100V | 20 | 3.8 | [5 empty] | 0.082916 | 115.8
50 | 100V | 50.0002 | 7 | [5 empty] | 0.061644 | 115.8
100 | 100V | 100.0004 | 6 |  |  |  | 2 |  | 0.042164 | 115.8
-10 | 100V | -10 | 3.8 | [5 empty] | -0.16415 | 115.8
-20 | 100V | -20.0001 | 3.8 | [5 empty] | -0.10607 | 115.8
-50 | 100V | -50.0002 | 7 | [5 empty] | -0.094398 | 115.8
-100 | 100V | -100.0004 | 6 |  |  |  | 2 |  | -0.025495 | 115.8
-100 | 1000V | -100.0004 | 6 | [5 empty] | -0.16583 | 115.8
-200 | 1000V | -200.0005 | 5.5 | [5 empty] | -0.058333 | 115.8
-500 | 1000V | -500.0026 | 8.2 | [5 empty] | -0.033333 | 115.8
-1000 | 1000V | -1000.0049 | 7.6 |  |  |  | 2 |  | -0.019365 | 115.8
100 | 1000V | 100.0002 | 6 | [5 empty] | 0.13944 | 115.8
200 | 1000V | 200.0004 | 5.5 | [5 empty] | 0.058333 | 115.8
500 | 1000V | 500.0022 | 8.2 | [5 empty] | 0.021082 | 115.8
1000 | 1000V | 1000.0046 | 7.6 |  |  |  | 2 |  | 0.017873 | 115.8
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99961 | 0.9996 | 85 |  |  |  | 20 |  | 0.25396 | 115.8
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.14913 | 115.8
10 Ω | 10.0004 | 10.0003 | 26 |  |  |  | 5 |  | 0.30595 | 115.8
19 Ω | 19.0039 | 19.0037 | 24 |  |  |  | 5 |  | 0.12372 | 115.8
100 Ω | 100.0053 | 100.0044 | 15 |  |  |  | 3 |  | 0.049998 | 115.8
190 Ω | 190.0032 | 190.0017 | 15 |  |  |  | 3 |  | 0.027739 | 115.8
1 kΩ | 999.9469 | 999.9364 | 11 |  |  |  | 3 |  | 0.066671 | 115.8
1.9 kΩ | 1899.9106 | 1899.8898 | 11 |  |  |  | 3 |  | 0.037218 | 115.8
10 kΩ | 10000.453 | 10000.379 | 9 |  |  |  | 3 |  | 0.06009 | 115.8
19 kΩ | 18999.722 | 18999.574 | 9 |  |  |  | 3 |  | 0.041145 | 115.8
100 kΩ | 99998.75 | 99998.03 | 11 |  |  |  | 5 |  | 0.052706 | 115.8
190 kΩ | 189994.66 | 189993.35 | 11 |  |  |  | 5 |  | 0.058846 | 115.8
1 MΩ | 999908.7 | 999911.4 | 16 |  |  |  | 8 |  | 0.23156 | 115.8
1.9 MΩ | 1899855.1 | 1899853.4 | 17 |  |  |  | 8 |  | 0.02321 | 115.8
10MΩ | 9998564 | 9999569 | 33 |  |  |  | 12 |  | 0.16415 | 115.8
19 MΩ | 18999500 | 19000526 | 43 |  |  |  | 12 |  | 0.22449 | 115.8
100MΩ(2W) | 100009390 | 100122640 | 110 |  |  |  | 180 |  | 1.3635 | 115.8
100MΩ(HIV 2W) | 100009390 | 100003500 | 110 |  |  |  | 180 |  | 0.70708 | 115.8
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 1 Ω | 3.8e-06 | 5e-05 | [5 empty] | 1.9003e-07 | 1.611e+39
0 Ω | 10 Ω | 3.3e-05 | 5e-05 |  |  | N/A |  |  | 2.9345e-06 | 1.611e+39
0 Ω | 100 Ω | 6e-05 | 5e-05 |  |  | N/A |  |  | 3.3333e-06 | 1.611e+39
0 Ω | 1 KΩ | 0.0005 | 5e-05 |  |  | N/A |  |  | 5e-05 | 1.611e+39
0 Ω | 10 KΩ | 0.005 | 5e-05 |  |  | N/A |  |  | 0.00044096 | 1.611e+39
0 Ω | 100 KΩ | -0.01 | 5e-05 |  |  | N/A |  |  | 0 | 1.611e+39
0 Ω | 1 MΩ | -0.1 | 5e-05 |  |  | N/A |  |  | 0.05 | 1.611e+39
0 Ω | 10 MΩ | -1 | 5e-05 |  |  | N/A |  |  | 0 | 1.611e+39
0 Ω | 100 MΩ(2W) | -20 | 5e-05 |  |  | N/A |  |  | 5 | 1.611e+39
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 1 Ω | 0.41028 | 5e-05 | [5 empty] | 0.00011093 | 1.611e+39
0 Ω | 10 Ω | 0.41009 | 5e-05 |  |  | N/A |  |  | 6.3521e-05 | 1.611e+39
0 Ω | 100 Ω | 0.40989 | 5e-05 |  |  | N/A |  |  | 4.5031e-05 | 1.611e+39
0 Ω | 1 KΩ | 0.4107 | 5e-05 |  |  | N/A |  |  | 4.4096e-05 | 1.611e+39
0 Ω | 10 KΩ | 0.419 | 5e-05 |  |  | N/A |  |  | 0.0005 | 1.611e+39
0 Ω | 100 KΩ | 0.4 | 5e-05 |  |  | N/A |  |  | 0 | 1.611e+39
0 Ω | 1 MΩ | 0.4 | 5e-05 |  |  | N/A |  |  | 0 | 1.611e+39
0 Ω | 10 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.44096 | 1.611e+39
0 Ω | 100 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 7.0711 | 1.611e+39
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99989 | 77 |  |  |  | 10 |  | 16.4166 | 115.8
2VAC/1kHz | 10 VAC | 1.9999 | 73.5 |  |  |  | 10 |  | 7.4073 | 115.8
5VAC/1kHz | 10 VAC | 4.9999 | 84 |  |  |  | 10 |  | 2.6458 | 115.8
10VAC/1kHz | 10 VAC | 9.9998 | 77 |  |  |  | 10 |  | 1.856 | 115.8
12VAC/1kHz | 10 VAC | 11.9998 | 75.833 |  |  |  | 10 |  | 1.559 | 115.8
15VAC/1kHz | 10 VAC | 14.9998 | 74.667 |  |  |  | 10 |  | 1.2222 | 115.8
19VAC/1kHz | 10 VAC | 18.9998 | 73.684 |  |  |  | 10 |  | 1.0886 | 115.8
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009963 | 5500 |  |  |  | 20 |  | 190.735 | 115.8
0.001 V, 20 Hz |  | 0.0009973 | 5200 |  |  |  | 20 |  | 66.8472 | 115.8
0.001 V, 30 Hz |  | 0.0009974 | 5200 |  |  |  | 20 |  | 105.684 | 115.8
0.001 V, 40 Hz |  | 0.0009974 | 5100 |  |  |  | 20 |  | 70.895 | 115.8
0.001 V, 55 Hz |  | 0.0009975 | 5100 |  |  |  | 20 |  | 88.4127 | 115.8
0.001 V, 300 Hz |  | 0.0009975 | 5100 |  |  |  | 20 |  | 347.2784 | 115.8
0.001 V, 1 KHz |  | 0.0009976 | 5100 |  |  |  | 20 |  | 284.015 | 115.8
0.001 V, 10 KHz |  | 0.0009977 | 5100 |  |  |  | 20 |  | 220.9875 | 115.8
0.001 V, 20 KHz |  | 0.0009976 | 5100 |  |  |  | 20 |  | 105.6628 | 115.8
0.001 V, 30 KHz |  | 0.0009973 | 5340 |  |  |  | 20 |  | 50.1354 | 115.8
0.001 V, 50 KHz |  | 0.0009951 | 5340 |  |  |  | 20 |  | 71.0589 | 115.8
0.001 V, 100 KHz |  | 0.0009809 | 8800 |  |  |  | 50 |  | 67.9648 | 115.8
0.001 V, 200 KHz |  | 0.0009268 | 16100 |  |  |  | 100 |  | 56.8673 | 115.8
0.001 V, 300 KHz |  | 0.0008947 | 16100 |  |  |  | 100 |  | 79.0328 | 115.8
0.001 V, 500 KHz |  | 0.0008577 | 31500 |  |  |  | 200 |  | 122.8976 | 115.8
0.001 V, 1 MHz |  | 0.0007737 | 33000 |  |  |  | 300 |  | 131.0319 | 115.8
0.01 V, 10 Hz |  | 0.0099861 | 1100 |  |  |  | 20 |  | 136.112 | 115.8
0.01 V, 20 Hz |  | 0.0099952 | 800 |  |  |  | 20 |  | 10.1428 | 115.8
0.01 V, 30 Hz |  | 0.0099968 | 800 |  |  |  | 20 |  | 9.7214 | 115.8
0.01 V, 40 Hz |  | 0.0099973 | 700 |  |  |  | 20 |  | 8.8216 | 115.8
0.01 V, 55 Hz |  | 0.0099975 | 700 |  |  |  | 20 |  | 15.903 | 115.8
0.01 V, 300 Hz |  | 0.0099976 | 700 |  |  |  | 20 |  | 20.8883 | 115.8
0.01 V, 1 KHz |  | 0.0099976 | 700 |  |  |  | 20 |  | 32.0233 | 115.8
0.01 V, 10 KHz |  | 0.0099991 | 700 |  |  |  | 20 |  | 21.7965 | 115.8
0.01 V, 20 KHz |  | 0.0099993 | 700 |  |  |  | 20 |  | 13.2297 | 115.8
0.01 V, 30 KHz |  | 0.01 | 940 |  |  |  | 20 |  | 12.6926 | 115.8
0.01 V, 50 KHz |  | 0.010003 | 940 |  |  |  | 20 |  | 10.5381 | 115.8
0.01 V, 100 KHz |  | 0.010008 | 1600 |  |  |  | 50 |  | 9.9921 | 115.8
0.01 V, 200 KHz |  | 0.0099956 | 2600 |  |  |  | 100 |  | 5.2728 | 115.8
0.01 V, 300 KHz |  | 0.0099928 | 2600 |  |  |  | 100 |  | 7.823 | 115.8
0.01 V, 500 KHz |  | 0.0099888 | 4500 |  |  |  | 200 |  | 7.273 | 115.8
0.01 V, 1 MHz |  | 0.0097846 | 6000 |  |  |  | 300 |  | 7.2267 | 115.8
0.1 V, 10 Hz |  | 0.099876 | 660 |  |  |  | 20 |  | 143.972 | 115.8
0.1 V, 20 Hz |  | 0.099966 | 300 |  |  |  | 20 |  | 4.1581 | 115.8
0.1 V, 30 Hz |  | 0.099981 | 300 |  |  |  | 20 |  | 1.1182 | 115.8
0.1 V, 40 Hz |  | 0.099986 | 195 |  |  |  | 20 |  | 2.1216 | 115.8
0.1 V, 55 Hz |  | 0.099989 | 195 |  |  |  | 20 |  | 3.167 | 115.8
0.1 V, 300 Hz |  | 0.099993 | 195 |  |  |  | 20 |  | 2.8482 | 115.8
0.1 V, 1 KHz |  | 0.099992 | 195 |  |  |  | 20 |  | 3.6327 | 115.8
0.1 V, 10 KHz |  | 0.099992 | 195 |  |  |  | 20 |  | 2.646 | 115.8
0.1 V, 20 KHz |  | 0.099987 | 195 |  |  |  | 20 |  | 1.8785 | 115.8
0.1 V, 30 KHz |  | 0.099983 | 400 |  |  |  | 20 |  | 1.4532 | 115.8
0.1 V, 50 KHz |  | 0.099983 | 400 |  |  |  | 20 |  | 1.0002 | 115.8
0.1 V, 100 KHz |  | 0.099982 | 1050 |  |  |  | 50 |  | 0.972 | 115.8
0.1 V, 200 KHz |  | 0.099961 | 1240 |  |  |  | 100 |  | 0.66692 | 115.8
0.1 V, 300 KHz |  | 0.099939 | 1240 |  |  |  | 100 |  | 0.92853 | 115.8
0.1 V, 500 KHz |  | 0.099913 | 1900 |  |  |  | 200 |  | 1.4542 | 115.8
0.1 V, 1 MHz |  | 0.099666 | 4000 |  |  |  | 300 |  | 1.6723 | 115.8
1 V, 10 Hz |  | 0.99868 | 600 |  |  |  | 10 |  | 131.5223 | 115.8
1 V, 20 Hz |  | 0.99972 | 180 |  |  |  | 10 |  | 3.4571 | 115.8
1 V, 30 Hz |  | 0.99987 | 180 |  |  |  | 10 |  | 2.1216 | 115.8
1 V, 40 Hz |  | 0.99992 | 77 |  |  |  | 10 |  | 2.186 | 115.8
1 V, 55 Hz |  | 0.99994 | 77 |  |  |  | 10 |  | 1.2694 | 115.8
1 V, 300 Hz |  | 0.99998 | 77 |  |  |  | 10 |  | 0.97184 | 115.8
1 V, 1 KHz |  | 0.99998 | 77 |  |  |  | 10 |  | 1.716 | 115.8
1 V, 10 KHz |  | 0.99998 | 77 |  |  |  | 10 |  | 1.4142 | 115.8
1 V, 20 KHz |  | 0.99997 | 77 |  |  |  | 10 |  | 0.92799 | 115.8
1 V, 30 KHz |  | 0.99997 | 140 |  |  |  | 10 |  | 1.1304 | 115.8
1 V, 50 KHz |  | 1 | 140 |  |  |  | 20 |  | 1 | 115.8
1 V, 100 KHz |  | 1.0001 | 310 |  |  |  | 30 |  | 0.99994 | 115.8
1 V, 200 KHz |  | 1.0001 | 550 |  |  |  | 70 |  | 0.78169 | 115.8
1 V, 300 KHz |  | 1.0001 | 550 |  |  |  | 70 |  | 0.70703 | 115.8
1 V, 500 KHz |  | 1.0009 | 1400 |  |  |  | 100 |  | 1.4987 | 115.8
1 V, 1 MHz |  | 1.0073 | 3000 |  |  |  | 200 |  | 2.7538 | 115.8
10 V, 10 Hz |  | 9.9869 | 600 |  |  |  | 10 |  | 142.1249 | 115.8
10 V, 20 Hz |  | 9.9973 | 180 |  |  |  | 10 |  | 3.9098 | 115.8
10 V, 30 Hz |  | 9.9988 | 180 |  |  |  | 10 |  | 2.4498 | 115.8
10 V, 40 Hz |  | 9.9993 | 77 |  |  |  | 10 |  | 1.2248 | 115.8
10 V, 55 Hz |  | 9.9996 | 77 |  |  |  | 10 |  | 2.5001 | 115.8
10 V, 300 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.5811 | 115.8
10 V, 1 KHz |  | 9.9999 | 77 |  |  |  | 10 |  | 2.8038 | 115.8
10 V, 10 KHz |  | 10.0002 | 77 |  |  |  | 10 |  | 1.5 | 115.8
10 V, 20 KHz |  | 10.0008 | 77 |  |  |  | 10 |  | 1.2246 | 115.8
10 V, 30 KHz |  | 10.0014 | 140 |  |  |  | 10 |  | 1.1665 | 115.8
10 V, 50 KHz |  | 10.0021 | 140 |  |  |  | 20 |  | 1.1664 | 115.8
10 V, 100 KHz |  | 10.0024 | 270 |  |  |  | 30 |  | 1.7634 | 115.8
10 V, 200 KHz |  | 10.0065 | 670 |  |  |  | 70 |  | 0.70665 | 115.8
10 V, 300 KHz |  | 10.0116 | 670 |  |  |  | 70 |  | 1.4126 | 115.8
10 V, 500 KHz |  | 10.0592 | 1700 |  |  |  | 100 |  | 1.0865 | 115.8
10 V, 1 MHz |  | 9.9958 | 3500 |  |  |  | 200 |  | 6.7852 | 115.8
19 V, 10 Hz |  | 18.9802 | 552.63 |  |  |  | 10 |  | 133.6273 | 115.8
19 V, 20 Hz |  | 18.996 | 165.79 |  |  |  | 10 |  | 2.3395 | 115.8
19 V, 30 Hz |  | 18.9983 | 165.79 |  |  |  | 10 |  | 1.9947 | 115.8
19 V, 40 Hz |  | 18.999 | 73.684 |  |  |  | 10 |  | 2.138 | 115.8
19 V, 55 Hz |  | 18.9994 | 73.684 |  |  |  | 10 |  | 1.761 | 115.8
19 V, 300 Hz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 0.76472 | 115.8
19 V, 1 KHz |  | 18.9998 | 73.684 |  |  |  | 10 |  | 1.6175 | 115.8
19 V, 10 KHz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 0.87717 | 115.8
19 V, 20 KHz |  | 19.0014 | 73.684 |  |  |  | 10 |  | 0.5949 | 115.8
19 V, 30 KHz |  | 19.0026 | 130.53 |  |  |  | 10 |  | 0.74937 | 115.8
19 V, 50 KHz |  | 19.0037 | 130.53 |  |  |  | 20 |  | 0.74933 | 115.8
19 V, 100 KHz |  | 19.0035 | 251.05 |  |  |  | 30 |  | 0.43852 | 115.8
19 V, 200 KHz |  | 19.0097 | 589.47 |  |  |  | 70 |  | 0.78907 | 115.8
19 V, 300 KHz |  | 19.0169 | 589.47 |  |  |  | 70 |  | 0.59441 | 115.8
19 V, 500 KHz |  | 19.1018 | 1463.2 |  |  |  | 100 |  | 1.2823 | 115.8
19 V, 1 MHz |  | 18.955 | 3073.7 |  |  |  | 200 |  | 6.2836 | 115.8
100 V, 10 Hz |  | 99.8778 | 600 |  |  |  | 10 |  | 128.4006 | 115.8
100 V, 20 Hz |  | 99.9737 | 180 |  |  |  | 10 |  | 3.6752 | 115.8
100 V, 30 Hz |  | 99.9887 | 180 |  |  |  | 10 |  | 1.764 | 115.8
100 V, 40 Hz |  | 99.9938 | 85 |  |  |  | 10 |  | 1.2694 | 115.8
100 V, 55 Hz |  | 99.9968 | 85 |  |  |  | 10 |  | 2.1279 | 115.8
100 V, 300 Hz |  | 100 | 85 |  |  |  | 10 |  | 1.4814 | 115.8
100 V, 1 KHz |  | 99.9996 | 85 |  |  |  | 10 |  | 1.7873 | 115.8
100 V, 10 KHz |  | 100.0019 | 85 |  |  |  | 10 |  | 1.3229 | 115.8
100 V, 20 KHz |  | 100.009 | 85 |  |  |  | 10 |  | 0.86595 | 115.8
100 V, 30 KHz |  | 100.0154 | 240 |  |  |  | 10 |  | 1.2017 | 115.8
100 V, 50 KHz |  | 100.023 | 240 |  |  |  | 20 |  | 0.83314 | 115.8
100 V, 100 KHz |  | 100.0281 | 600 |  |  |  | 30 |  | 3.9359 | 115.8
100 V, 200 KHz |  | 100.113 | 2600 |  |  |  | 50 |  | 15.3702 | 115.8
1000 V, 55 Hz |  | 999.957 | 79 |  |  |  | 15 |  | 2.4553 | 115.8
1000 V, 300 Hz |  | 1000.006 | 79 |  |  |  | 15 |  | 1.0138 | 115.8
700 V, 1 KHz |  | 700.011 | 80.714 |  |  |  | 50 |  | 1.7169 | 115.8
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.9983e-07 | 10046.7079 | [5 empty] | 18.2605 | 115.8
10 µADC | 1e-05 | 1e-05 | 1044.9582 | [5 empty] | 4.1128 | 115.8
50 µADC | 5e-05 | 5.0003e-05 | 245 | [5 empty] | 0.34154 | 115.8
100 µADC | 0.0001 | 0.00010001 | 145 |  |  |  | 7 |  | 0.14141 | 115.8
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | -0.14141 | 115.8
-50 µADC | -5e-05 | -5.0003e-05 | 245 | [5 empty] | -0.29998 | 115.8
-10 µADC | -1e-05 | -1.0001e-05 | -1044.9122 | [5 empty] | -2.0614 | 115.8
0.1 mADC | 0.0001 | 0.00010001 | 145 | [5 empty] | 0.95737 | 115.8
0.5 mADC | 0.0005 | 0.00050002 | 65 | [5 empty] | 0.16329 | 115.8
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0.12909 | 115.8
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | -0.049998 | 115.8
-0.5 mADC | -0.0005 | -0.00050002 | 65 | [5 empty] | -0.16329 | 115.8
1 mADC | 0.001 | 0.001 | 55 | [5 empty] | 1.2909 | 115.8
5 mADC | 0.005 | 0.0050002 | 65 | [5 empty] | 0.36513 | 115.8
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0.22173 | 115.8
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | -0.095739 | 115.8
-5 mADC | -0.005 | -0.0050002 | 65 | [5 empty] | -0.36513 | 115.8
10 mADC | 0.01 | 0.01 | 55 | [5 empty] | 1.291 | 115.8
50 mADC | 0.05 | 0.050005 | 75 | [5 empty] | 0.25818 | 115.8
100 mADC | 0.1 | 0.10001 | 65 |  |  |  | 7 |  | 0.049995 | 115.8
-100 mADC | -0.1 | -0.10001 | 65 |  |  |  | 7 |  | -0.12909 | 115.8
-50 mADC | -0.05 | -0.050005 | 75 | [5 empty] | -0.34153 | 115.8
0.1 ADC | 0.1 | 0.099939 | 65 | [5 empty] | 0.817 | 115.8
0.5 ADC | 0.5 | 0.4998 | 135 | [5 empty] | 0.16336 | 115.8
-0.5 ADC | -0.5 | -0.49989 | 135 | [5 empty] | -0.16334 | 115.8
-1.0 ADC | -1 | -0.99972 | 105 |  |  |  | 20 |  | -0.095769 | 115.8
1.0 ADC | 1 | 0.99964 | 105 |  |  |  | 20 |  | 0.14147 | 115.8
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.9907e-06 | 3653.3794 | [5 empty] | 214.4412 | 115.8
10 µA, 20 Hz | 1e-05 | 9.9922e-06 | 2852.2105 | [5 empty] | 9.5817 | 115.8
10 µA, 30 Hz | 1e-05 | 9.9923e-06 | 2852.1962 | [5 empty] | 18.2715 | 115.8
10 µA, 40 Hz | 1e-05 | 9.9916e-06 | 2121.7929 | [5 empty] | 20.633 | 115.8
10 µA 55 Hz | 1e-05 | 9.9922e-06 | 2121.6443 | [5 empty] | 24.5139 | 115.8
10 µA, 300 Hz | 1e-05 | 9.9921e-06 | 2121.6761 |  |  | -40 |  |  | 0 | 115.8
10 µA, 1 KHz | 1e-05 | 9.9917e-06 | 2121.7504 | [5 empty] | 5.7783 | 115.8
10 µA, 5 KHz | 1e-05 | 9.9822e-06 | 5509.8351 | [5 empty] | 9.5914 | 115.8
10 µA, 10 KHz | 1e-05 | 9.953e-06 | 11554.2472 | [5 empty] | 19.019 | 115.8
100 µA, 10 Hz | 0.0001 | 0.0001 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 145.4628 | 115.8
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 0.95734 | 115.8
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 2.1601 | 115.8
100 µA, 40 Hz | 0.0001 | 0.0001 | 320 |  |  |  | 50 |  | 1.2583 | 115.8
100 µA 55 Hz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 5.8522 | 115.8
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 0.95737 | 115.8
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 0.95742 | 115.8
100 µA, 5 KHz | 0.0001 | 9.9914e-05 | 0.1 |  |  |  | 100 |  | 1.7093 | 115.8
100 µA, 10 KHz | 0.0001 | 9.9652e-05 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 1.2955 | 115.8
1mA, 10 Hz | 0.001 | 0.00099996 | 700 |  |  |  | 40 |  | 153.7029 | 115.8
1 mA, 20 Hz | 0.001 | 0.001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.4999 | 115.8
1 mA, 30 Hz | 0.001 | 0.001 | 390 |  |  |  | 40 |  | 1.2583 | 115.8
1 mA, 40 Hz | 0.001 | 0.00099996 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.6301 | 115.8
1 mA, 55 Hz | 0.001 | 0.00099999 | 160 |  |  |  | 40 |  | 1.5 | 115.8
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95741 | 115.8
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 2.2173 | 115.8
1 mA, 5 KHz | 0.001 | 0.00099996 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 2.2174 | 115.8
1 mA, 10 KHz | 0.001 | 0.00099973 | 0.25 |  |  |  | 200 |  | 1.0003 | 115.8
10 mA, 10 Hz | 0.01 | 0.0099994 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 177.7756 | 115.8
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 2.4494 | 115.8
10 mA, 30 Hz | 0.01 | 0.01 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.8257 | 115.8
10 mA, 40 Hz | 0.01 | 0.0099996 | 160 |  |  |  | 40 |  | 2.2174 | 115.8
10 mA, 55 Hz | 0.01 | 0.0099999 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.291 | 115.8
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 0.95741 | 115.8
10 mA, 1 KHz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.5 | 115.8
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 |  |  |  | 70 |  | 1.7319 | 115.8
10 mA, 10 KHz | 0.01 | 0.010001 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 1.4141 | 115.8
100 mA, 10 Hz | 0.1 | 0.099991 | 700 |  |  |  | 40 |  | 133.8336 | 115.8
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 3.5936 | 115.8
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 1.4999 | 115.8
100 mA, 40 Hz | 0.1 | 0.1 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.893 | 115.8
100 mA, 55 Hz | 0.1 | 0.1 | 160 |  |  |  | 40 |  | 1.4142 | 115.8
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49996 | 115.8
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 3.1089 | 115.8
100 mA, 5 KHz | 0.1 | 0.10002 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 2.0611 | 115.8
100 mA, 10 KHz | 0.1 | 0.10005 | 0.25 |  |  |  | 200 |  | 2.2162 | 115.8
1 A, 10 Hz | 1 | 0.99965 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 135.5123 | 115.8
1 A, 20 Hz | 1 | 0.99969 | -1 |  |  |  | 40 |  | 2.987 | 115.8
1 A, 30 Hz | 1 | 0.99967 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.2181 | 115.8
1 A, 40 Hz | 1 | 0.99959 | 640 |  |  |  | 40 |  | 3.5014 | 115.8
1 A, 55 Hz | 1 | 0.99963 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 13.3059 | 115.8
1 A, 300 Hz | 1 | 0.99966 | 640 |  |  |  | 40 |  | 2.3102 | 115.8
1 A, 1 KHz | 1 | 0.9997 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 4.7972 | 115.8
1 A, 5 KHz | 1 | 0.99973 | 800 |  |  |  | 70 |  | 2.3811 | 115.8
1 A, 10 KHz | 1 | 0.99932 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 2.2189 | 115.8
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.01 | 278.29 | 185.95 | 166.12 | 33.72 | 27.74